## Slide 1: NinjaNIRS 2022 Firmware Schematic and Data Format Specification
Version 2023-8-17
Bernhard Zimmermann – bzim@bu.edu

## Slide 2: Control Card Block Diagram
PSRAM
FPGA
RPi Zero
RP2040
FTDI USB
PSRAM
Detector Backplanes
Source Backplane
Detector Backplanes
Detector Backplanes
Detector Backplanes
UART + SWD + 1 GPIO
9 GPIO (incl RPi 0 SPI)
Spare Con.
6 GPIO
12 GPIO
UART with RTS/CTS
7 GPIO
Spare Con2
4 GPIO
Aux ADCs
2x6 GPIO
15 GPIO UART with CTS/RTS or parallel data
2 GPIO (incl I2C)
4 Backplane Connectors
Per Connector 17 GPIO
(5 common + 4x3 direct)
1 Backplane Connector
21 GPIO
(3 rows + 7x2 columns + 3 power select + 1 spare)
Aux DigIO
Remote Dig In
6 GPIO
2 GPIO
4 GPIO

## Slide 3: FPGA Firmware Block Diagram
UART Rx
UART Tx
18 x 1024
Acquisition Sequence RAM B
27 x 1024
State Sequence RAM A
Source Backplane
8 rows
8 columns
3 power select
Program
Counter A
Program
Counter B
1 Trigger
1 Trigger Status
1 End Cycle
16 Board Select
+1
FTDI USB
Host PC
USB
Config
FSM
Acq
FSM/MUX
UART Rx
Aux ADC Interface
Detector Backplanes
FIFO
UART with RTS/CTS
Status Registers
FPGA
readback
Clock Divider
+1

## Slide 4: RAM Data Formats
RAM A: State Sequence
End: 1 if last state in sequence: will jump to state 0 next; 0 otherwise: will jump to state i+1 next
X Reserved: reserved for later use
Status: ‘1’ send status packets instead of regular light level packets; ‘0’ send regular packets
Power: select LED power level. 0b001: high power, 0b010: mid power, 0b100: low power, 0b000: off (dark state)
Column select: Select column of LED to be illuminated. Even numbers correspond to XXX nm, odd numbers to YYY nm. This is selecting source boards.
Row select: Select row of LED to be illuminated. This is switching between LEDs in a source board
 | 26 | 25 | 24 | 23 | 22 | 21 | 20 | 19 | 18 | 17 | 16 | 15 | 14 | 13 | 12 | 11 | 10 | 9 | 8 | 7 | 6 | 5 | 4 | 3 | 2 | 1 | 0
state 0 | end | X reserved |  |  |  | stat | power | stat | power | column select |  | column select | data source |  | packet count |  |  | Lsr id |  |  |  |  |  | X | row select |  | 
⁞ |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
S 1023 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
RAM B: Acquisition Sequence
End: 1 if last state in sequence: will jump to state 0 next; 0 otherwise: will jump to state i+1 next
X Reserved: reserved for later use
Acq Trg: Transition form ‘0’ to ‘1’ will trigger the acquisition of one sample in the detector boards
End Cyc: Transition form ‘0’ to ‘1’ will cause the detector boards to move the current measurements to the transmit buffer, and reset the sample sums to 0
PC Tx Trg: Program counter A transmit trigger. Transition form ‘0’ to ‘1’ will cause the current program counter A (state) to be transmitted
Aux Tx Trg: Auxillary transmit trigger. Transition form ‘0’ to ‘1’ will cause the current auxiliary ADC values to be transmitted.
DetB Rx En: Enable reception of data from the detector board selected by UART Rx Mux. Sets CTS for the selected board.
UART Rx Mux: Selects from which detector board to receive data.
 | 17 | 16 | 15 | 14 | 13 | 12 | 11 | 10 | 9 | 8 | 7 | 6 | 5 | 4 | 3 | 2 | 1 | 0
State 0 | End | X | X | X | X | X | X | X | Acq Trg | End cyc | PC Tx Trg | Aux Tx Trg | DetB Rx En | X | UART Rx Mux |  |  | 

## Slide 5: Status Register Format
Status Register
Clk Div: Clock divider setting. RAM B state period = ClkDiv*8/96MHz
Rst Ram: Reset the PSRAM chip used in the external FIFO.
Rst PC A: Reset program counter A to 0
Rst DetB: Reset the microcontroller on the detector boards by pulling the RUN pins low. Each bit corresponds to one backplane.
Run: ‘1’: the clock divider and program counters A and B will run. ‘0’: Divider continues to run until PC B and the divider reach Zero, and then stay there. LED power will be set to 0 immediately.
 | 23 | 22 | 21 | 20 | 19 | 18 | 17 | 16 | 15 | 14 | 13 | 12 | 11 | 10 | 9 | 8 | 7 | 6 | 5 | 4 | 3 | 2 | 1 | 0
 | VN22Clk en | VN3P4 en | VN22 en | V9P0 en | V5P1src en | V5P1RPi en | V5P1B23 en | V5P1B01 en | Clk Div |  |  |  |  |  |  |  | X | Rst Ram | Rst PC A | Rst DetB |  |  | Rst DetB | Run

## Slide 6: Command Format
Command sequence
Each command starts with a header byte (‘255’), then 2 address bytes, then 4 command payload bytes.
Addr: Address of data to be written or read. (see below)
Rdat: Register or RAM data to be written. Write all ‘0’ when reading a register / RAM location.
 | byte[0] | byte[1] | byte[2] | byte[3] | byte[4] | byte[5] | byte[6]
 | 255 | addr[0] | addr[1] | rdat[0] | rdat[1] | rdat[2] | rdat[3]
Command address format
 | 15 | 14 | 13 | 12 | 11 | 10 | 9 | 8 | 7 | 6 | 5 | 4 | 3 | 2 | 1 | 0
 | addr[1] |  |  |  |  |  |  |  | addr[0] |  |  |  |  |  |  | 
 | r̄/w | addr | X | X | X | X | Acq Trg | End cyc | PC Tx Trg | Aux Tx Trg | DetB Rx En | X | UART Rx Mux |  |  | 
r̄/w: ‘1’: write to the location specified by ‘addr’. ‘0’: read from location ‘addr’. Read is only save if Run = ‘0’.
Addr: Address to be read or written.
0b001 0000 0000 0000 through 0b 001 0011 1111 1111 : RAM A
0b010 0000 0000 0000 through 0b 001 0011 1111 1111 : RAM B
0b100 0000 0000 0001 : Status Register

## Slide 7: Output Data Packet Format
Note: The order of the individual data blocks below is determined by the contents of RAM A. The order shown below is a typical example as also implemented in the MATLAB helper functions.
Output data sequence
n: number of plugged in Detector Adapter Boards
See following slides for details on the individual packets.
 | byte[0..2] | byte[(0..7)+3] | byte[(0..27)+11] | … | byte[(0..27)+(n-1)*28+11] | byte[(0..17)+n*28+11]
 | PC A packet | Auxiliary pkt | Det adapter pkt | … | Det adapter pkt | Accelerometer pkt

## Slide 8: Program Counter A Packet Format
Packet generated by the FPGA, usually placed at the beginning of a sample/RAM B cycle to help framing the data. See also PC TX Trg signal in RAM B data format slide.
Output data sequence
PC A [1]: high byte of program counter A
PC A [0]: low byte of program counter A
 | byte[0] | byte[1] | byte[2]
 | 254 | PC A [0] | PC A [1]

## Slide 9: Auxiliary Packet Format
Data packet generated by the FPGA. Contains the values of the two on-board auxiliary analog to digital converters (ADC). The ADCs are sampled and averaged at the same time and in the same manner as the main ADCs on the detector boards.
The packet also contains the status of the remote trigger signals and the Aux Digital inputs. These signals are sampled a few clock cycles after the rising edge of the “Aux Trg Tx” signal (see RAM B description).
Output data sequence
Dig In: Remote trigger and aux digital inputs sampled at time of “Aux Trg Tx” signal.
Bits (0..3): remote trigger signals.
Bits (4..5): aux digital inputs.
ADCx: 24 bit sum of all conversions since last “Aux Trg Tx” signal. [2] is the high byte (little endian). Unsigned integer.
 | byte[0] | byte[1] | byte[2] | byte[3] | byte[4] | byte[5] | byte[6] | byte[7]
 | 250 | Dig In | ADC0[0] | ADC0[1] | ADC0[2] | ADC1[0] | ADC1[1] | ADC1[2]

## Slide 10: Detector Adapter Packet Format
Data packet from the RP2040 microcontroller on a detector adapter board. Contains the main light intensity data.
Output data sequence
Smp cnt: Count of transmitted samples (modulo 256) since last restart of the detector adapter card.
ADCx: 24 bit sum of all conversions since last ‘end cycle’ signal. [2] is the high byte (little endian). Signed integer / two’s complement.
 | byte[0] | byte[1] | byte[2] | byte[3] | byte[4] | byte[5] | … | byte[24] | byte[25] | byte[26] | byte[27]
 | 253 | 252 | Smp cnt | ADC0[0] | ADC0[1] | ADC0[2] |  | ADC7[0] | ADC7[1] | ADC7[2] | 0

## Slide 11: Detector Adapter Status Packet Format
Status data packet from the RP2040 microcontroller on a detector adapter board.
Output data sequence
Smp cnt: Count of transmitted samples (modulo 256) since last restart of the microcontroller.
UID: 64 bit unique ID from flash attached to microcontroller
FW: Firmware version.
Temp: Microcontroller internal temperature.
Conn flag: accelerometer connected flags (acc_connected_before_run<<1) || acc_connected_curr)
 | byte[0] | byte[1] | byte[2] | byte[3] | … | byte[10] | byte[11] | byte[12] | byte[13] | byte[14] | byte[15] | byte[16] | byte[17]
 | 237 | 236 | Smp cnt | UID[0] |  | UID[7] | FWv[0] | FWv[1] | Temp[0] | Temp[1] | 0 | 0 | 0

## Slide 12: Accelerometer Packet Format
Data packet from the RP2040 microcontroller on the main control board. Contains the data read out from an IMU attached to the QWIIC port.
Output data sequence
Smp cnt: Count of transmitted samples (modulo 256) since last restart of the microcontroller.
Datax: 16 bit data words from the IMU, obtained right after last “End cyc” signal. [1] is the high byte (little endian). Signed integer / two’s complement.
Data0: Temperature
Data1: Gyro X Data2: Gyro Y Data3: Gyro Z Data4: Accelerometer X Data5: Accelerometer Y Data6: Accelerometer Z
 | byte[0] | byte[1] | byte[2] | byte[3] | byte[4] | … | byte[15] | byte[16] | byte[17]
 | 249 | 248 | Smp cnt | Data0[0] | Data0[1] |  | Data6[0] | Data6[1] | 0

## Slide 13: Accelerometer Status Packet Format
Status data packet from the RP2040 microcontroller on the main control board. Contains the controller temperature and board input voltage.
Output data sequence
Smp cnt: Count of transmitted samples (modulo 256) since last restart of the microcontroller.
UID: 64 bit unique ID from flash attached to microcontroller
FW: Firmware version.
Temp: Microcontroller internal temperature.
Vin: Control board input voltage.
Conn flag: accelerometer connected flags (acc_connected_before_run<<1) || acc_connected_curr)
 | byte[0] | byte[1] | byte[2] | byte[3] | … | byte[10] | byte[11] | byte[12] | byte[13] | byte[14] | byte[15] | byte[16] | byte[17] | byte[18] | byte[19]
 | 237 | 236 | Smp cnt | UID[0] |  | UID[7] | FWv[0] | FWv[1] | Temp[0] | Temp[1] | Vin[0] | Vin[1] | Conn flag | 0 | 0

## Slide 14: Summary of Packet Headers
Summary of packet headers from the different data sources.
Packet Type | Header
Command from PC to FPGA | [255]
Reply to command (FPGA => PC) | [240]
Program Counter A Packet | [254]
Auxiliary Data Packet | [250]
Detector Adapter Data Packet | [253 252]
Detector Adapter Status Packet | [239 238]
IMU Controller Data Packet | [249 248]
IMU Controller Status Packet | [237 236]